UDecide

Verification test protocol 08   
DOCUMENT REFERENCE NUMBER: UD-VeTPr08  
  
DATE: 05-05-2021  
 REVISION NUMBER: 2.0

*AUTHOR:* Emma Elbo

*REVIEWER:* Sofie Bjørn, Sigrid Stang & Amalie Koch

SCOPE: This document contains the verification test protocol for UD-SRS-15.

REFERENCES: UD-SRS and UD-SRS-traceability

REVISION HISTORY:

|  |  |  |  |
| --- | --- | --- | --- |
| **Revision** | **Revised by** | **Revision date** | **Description of changes** |
| 1.0 | Sigrid Stang | 03-05-2021 | First version of verification test protocol number 08. |
| 2.0 | Emma Elbo | 05-05-2021 | Second version of verification test protocol number 08. Acceptance criteria, prerequisites and procedure are added. |
| 2.1 | Emma Elbo | 17-05-2021 | The acceptance criterion was changed in accordance with the UD-SRS-15 requirement. |
|  |  |  |  |

APPROVAL:

|  |  |  |  |
| --- | --- | --- | --- |
|  | *Name and job function* | *Signature* | *Date* |
| *Author:* |  | | |
| *Reviewer:* |  | | |
| *Independent reviewer:* |  | | |

|  |  |
| --- | --- |
| **SRS ID** | **Functional requirement** |
| UD-SRS-15 | Patient preferences regarding electrode type, shall be possible to input in the Udecide system |

|  |  |
| --- | --- |
| **TestID** | VeTPr08 |
| **Acceptance criteria** | There must be two input options on the user interface, where only one of them can be checked. The two input options shall be surface and percutaneous. |
| **Prerequisites** | Verification test protocol 08 can only be used to testing if the verification test 01 is well conducted:  The four different functionalities “login”, “search patient”, “estimate effectiveness score” and “report effect” are working on their respective interfaces. |
| **Procedure** | 1. Login to the UDecide system and navigate to the search patient view. 2. Find or create a patient in the UDecide system. 3. Click on to the questionnaire view and complete the questionnaire for the patient. 4. Click on to the recommended treatment view, where three recommended treatment and a check box for *choose electrode preference* are visualized. 5. Tick off “surface” or “percutaneous” electrode preference in the checkbox. |