UDecide

Validation test report 05   
DOCUMENT REFERENCE NUMBER: UD-VaTR05  
  
DATE: 18-05-2021  
 REVISION NUMBER: 1.0

*AUTHOR:* Emma Elbo

*REVIEWER:* Sofie Bjørn, Sigrid Stang & Amalie Koch

SCOPE: This document contains the validation test report for UD-CRS-11.

REFERENCES: UD-CRS.

REVISION HISTORY:

|  |  |  |  |
| --- | --- | --- | --- |
| **Revision** | **Revised by** | **Revision date** | **Description of changes** |
| 1.0 | Emma Elbo | 18-05-2021 | First version of validation test report number 05. |
|  |  |  |  |

APPROVAL:

|  |  |  |  |
| --- | --- | --- | --- |
|  | *Name and job function* | *Signature* | *Date* |
| *Author:* |  | | |
| *Reviewer:* |  | | |
| *Independent reviewer:* |  | | |

|  |  |
| --- | --- |
| **CRS ID** | **Requirement** |
| UD-CRS-11 | The output of the UDecide system shall weight the output, so that a minimal stimulation-time relative to high effect is weighted as the most beneficial stimulation paradigm. |

|  |  |
| --- | --- |
| **Test ID** | **VaTR05** |
| **Test protocol** | VaTP05 |
| **Results** | Only the recommended treatment with the estimated highest effect is highlighted. |
| **Acceptance criteria** | On the recommended treatment page, one of the three recommended treatment for the patients shall be highlighted. The highlighted treatment shall be the treatment with minimal stimulation-time relative to high effect is weighted as the most beneficial stimulation paradigm. |
| **Validated** | Not validated |