UDecide

Verification test report 06   
DOCUMENT REFERENCE NUMBER: UD-VeTR06  
  
DATE: 13-05-2021  
 REVISION NUMBER: 2.0

*AUTHOR:* Emma Elbo

*REVIEWER:* Sofie Bjørn, Sigrid Stang & Amalie Koch

SCOPE: This document contains the verification test for UD-SRS-11.

REFERENCES: UD-SRS and UD-SRS-traceability.

REVISION HISTORY:

|  |  |  |  |
| --- | --- | --- | --- |
| **Revision** | **Revised by** | **Revision date** | **Description of changes** |
| 1.0 | Emma Elbo | 06-05-2021 | First version of verification test report number 06. |
| 2.0 | Emma Elbo | 13-05-2021 | Added test results to the report |
|  |  |  |  |

APPROVAL:

|  |  |  |  |
| --- | --- | --- | --- |
|  | *Name and job function* | *Signature* | *Date* |
| *Author:* |  | | |
| *Reviewer:* |  | | |
| *Independent reviewer:* |  | | |

|  |  |
| --- | --- |
| **SRS ID** | **Functional requirement** |
| UD-SRS-11 | Stimulation paradigm recommended in the output from the Udecide system shall be either: urge, time limited or constant, and the recommended session time shall be either: 60 seconds, 15 minutes, 30 minutes, 4 hours or constant. |

|  |  |
| --- | --- |
| **Test ID** | **VeTR06** |
| **Test protocol** | VeTP06 |
| **Results** | The UDecide system can recommend following treatment strategy:   * Timed limited stimulation: 4 hours (Intensity: low to high) * Timed limited stimulation: 30 minutes (Intensity: low to high) * Timed limited stimulation: 15 minutes (Intensity: low to high) * Constant stimulation: 24 hours/constant (Intensity: low to high) * Urge stimulation: 60 seconds (Intensity: low to high) |
| **Acceptance criteria** | The different stimulation paradigms and parameters are consistent. Urge is paired with 60 seconds. Time limited is paired with either 15 min, 30 min or 4 hrs. Continuous is paired with constant. |
| **Verified** | Verified |