## Ark1
SRS ID | Requirement | CRS# | Additional documentation for the requirement | RM# | VeTPr#
UD-SRS-01 | The Udecide system shall consist of four functionalities: login, search patient, estimate effectiveness score and report effect. | UD-CRS-03 | Use case
UD-SRS-02 | The UDecide system shall recommend and visualize three treatment strategies for the specialist | UD-CRS-01
UD-SRS-03 | The Udecide system shall be able to calculate an effectiveness score based on inputs regarding symptoms before and after treatment. | UD-CRS-06 UD-CRS-07
UD-SRS-04 | The Udecide system shall rank the recommended treatment strategies from best to least expected effect | UD-CRS-01
UD-SRS-05 | The treatment strategy with the best expected effect shall be highlighted | UD-CRS-03
UD-SRS-06 | Each recommended treatment strategy shall contain information about stimulation paradigm, parameters and effectiveness score to give the specialist the opportunity to get an overview. | UD-CRS-03
UD-SRS-07 | The interface shall convey cautionary notes about the treatment to the specialist. | UD-CRS-01
UD-SRS-08 | The recommended treatment with the least time of stimulation pr. day shall be highlighted in the recommendation. | UD-CRS-11 |  | UD-H2
UD-SRS-09 | The system shall be able to take inputs to the algorithm entered by the specialist. Those inputs shall be: gender, age, number of incontinence episodes, number of daily urinations, nocturia, number of urgency episodes, bladder capacity, detrusor overactivity and quality of life. | UD-CRS-05 | Fluid/urination schedule | UD-H3
UD-SRS-10 | The Udecide system shall have a range check on the data input from the specialist | UD-H3 |  | UD-H3
UD-SRS-11 | Stimulation paradigm recommended in the output from the Udecide system shall be either: urge, time limited or constant, and the recommended session time shall be either: 60 seconds, 15 minutes, 30 minutes, 4 hours or constant. | UD-CRS-08 | Ucon Sheet
UD-SRS-12 | The predetermined stimulation parameters defined by InnoCon Medical shall be shown on the interface. | UD-CRS-02 |  | UD-H1
UD-SRS-13 | Information on the actual use of UCon shall be conveyed through graphical illustrations | UD-CRS-03 UD-CRS-06 UD-CRS-07
UD-SRS-14 | The questionnaire and the subsequent recommendation shall be conveyed in medical terms | UD-CRS-03 UD-CRS-04
UD-SRS-15 | Patient preferences regarding electrode type, shall be possible to input in the Udecide system | UD-CRS-10
UD-SRS-16 | The patient profile interface shall include input fields for creating a new patient profile. Those fields shall be: CPR, name, age and gender. | UD-CRS-12
UD-SRS-17 | The Udecide system shall only grant access for validated users. | H6 | MDR 2017-745 chapter 2, article 17. Used to to secure protection against unauthorised acces. | UD-H6
UD-SRS-18 | The Udecide system shall be connected to a cloud-based database. | UD-CRS-09 UD-CRS-12 UD-CRS-13 | Post-market-surveillance
UD-SRS-19 | The UDecide system shall be able to update the algorithm with new data on the research field periodically. | UD-H4 | OBS! ER identisk med CRS-09 | UD-H4 UD-H5
UD-SRS-20 | The patient profile, questionnaire, chosen treatment and treatment effect shall be stored in the database. | UD-CRS-12 UD-CRS-13
UD-SRS-21 | The Udecide system shall be compatible with a computer | UD-CRS-04
UD-SRS-22 | The software shall be programmed in a platform independent language | UD-CRS-04
UD-SRS-23 | The Udecide system shall handle patient data in accordance to general data protection regulation (GDPR) | UD-CRS-12