## Ark1
SRS ID | Requirement | CRS# | Additional documentation for the requirement | RM# | VeTPr#
UD-SRS-01 | The interface shall recommend and visualize three treatment strategies for the specialist | UD-CRS-01
UD-SRS-02 | The interface shall rank the recommended treatment strategies from best to least expected effect | UD-CRS-01
UD-SRS-03 | The interface shall convey practical information about the treatment to the specialist. The practical information shall include a describtion of the electrode and typical usage. | UD-CRS-01
UD-SRS-04 | The predetermined parameters defined by InnoCon Medical shall be shown on the interface. | UD-CRS-02
UD-SRS-05 | The Udecide system shall consist of three functionalities: login, choose patient and estimate effectiveness score | UD-CRS-03 | Use case
UD-SRS-06 | The user interface "Estimate effectiveness score" shall contain information about stimulation paradigm, parameters and effectiveness score to give the specialist the opportunity to get an overview. | UD-CRS-03
UD-SRS-07 | The system shall be able to update the algorithm with new data on the research field periodically | UD-H4
UD-SRS-08 | The output shall be conveyed through graphical illustrations | UD-CRS-03 UD-CRS-06 UD-CRS-07
UD-SRS-09 | The information from the output shall be conveyed in medical terms | UD-CRS-03 UD-CRS-04
UD-SRS-10 | The treatment strategy with the best expected effect shall be highlighted | UD-CRS-03
UD-SRS-11 | The Udecide system shall be compatible with a computer | UD-CRS-04
UD-SRS-12 | The software shall be programmed in a platform independent language | UD-CRS-04
UD-SRS-13 | The system shall be able to take inputs to the algorithm entered by the specialist. Those inputs shall be: sex, age, OAB symptoms and data collected from clinical practice | UD-CRS-05 | Fluid/urination schedule
UD-SRS-14 | The output of the system shall represent pros and cons for each recommended treatment strategy | UD-CRS-06
UD-SRS-15 | Udecide shall have a range check on the data input from the specialist | UD-H3
UD-SRS-16 | The algorithm shall be based on data from UCon collected by InnoCon Medical | UD-CRS-02 UD-CRS-09
UD-SRS-17 | The Udecide system shall handle patient data in accordance to general data protection regulation (GDPR) | UD-CRS-12
UD-SRS-18 | Stimulation paradigm recommended in the output from the Udecide system shall be either: "Urge", "Time Limited" or "Constant". The output of stimulation parameter: "Intensity limit" shall be divided into 5 levels (level 1 = low, level 5 = high) and the output of stimulation parameter "Session time" shall be either: "60 second", "15 minutes", "30 minutes", "4 hours" or "constant". | UD-CRS-08 | Ucon Sheet
UD-SRS-19 | The Udecide system shall be connected to a database. | UD-CRS-09 UD-CRS-12 UD-CRS-13 | Post-market-surveillance
UD-SRS-20 | Patient preferencenses regarding electrode type, shall be an opportunity to input to the Udecide system | UD-CRS-10
UD-SRS-21 | The recommended treatment with the least time of stimulation pr. day shall be highlighted on the interface. | UD-CRS-11
UD-SRS-22 | The "Choose patient" interface shall include input fields for creating a new patient profile. Those field shall be: "CPR", "Name" and "Sex". The patient profile shall be stored in the database. | UD-CRS-12
UD-SRS-23 | The Udecide system shall validate the user access is granted | H6 | MDR 2017-745 chapter 2, article 17. Used to to secure protection against unauthorised acces.