## Ark1
Hazard-ID | Hazard | Cause of hazarduous situation | Sequence of events | Hazardouos situation | Harm | Object at risk | Initial risk assesment |  |  |  | Risk control | Residual risk estimation |  |  | Acceptance | CRS# SRS# | VeTP#
[7 empty] | Severity | Occurence | Risk (RPN) | Class |  | Severity | Occurence | Risk (RPN)
UD-H1 | Electric energy | UCon stimulates at a frequency too high | 1. Udecide recommends a frequency that is too high. 2. Clinician trusts the Udecide system and initiates treatment with the high frequency. | Patient is exposed to high frequency stimulation. | Nerve damage | Patient | 4 | 2 | 8 | Intolerable | Udecide shall not be able to recommend a higher frequency than 20 Hz. (UCon solely operates at 20 Hz) | 3 | 1 | 3 | Accepted | UD-CRS-02 UD-SRS-04
UD-H2 | Immonulogical, irritants | Use of UCon electrodes | 1. Udecide recommends a treatment with a specific type of electrode. 2. Patient has a skin reaction to the electrode. | Patient is exposed to material that can potentially irritate the skin. | Skin irritation | Patient | 2 | 5 | 10 | Intolerable | Udecide shall recommend treatment with the shortest stimulation time as possible. | 2 | 3 | 6 | Accepted | UD-CRS-11 UD-SRS-21
UD-H3 | Operational, use error | Input is wrong | 1. Udecide requests patient data to the algorithm. 2. Clinician inputs wrong data. | Patient is categorized wrong and thus receives a suboptimal treatment recommendation. | Suboptimal treatment | Patient | 1 | 5 | 5 | Broadly acceptable | a. Udecide has a range check for data input from the user. b. The interface shall show a warning to force the clinician to check the information input to the system twice. | 1 | 3 | 3 | Accepted | UD-SRS-13 UD-SRS-15
UD-H4 | Operational, function | Outdated algorithm | 1. There is new knowledge the research field, that has not been implemented in the system. 2. The clinician uses Udecide in an unupdated version. 3. The clinician recieves decision support that is outdated. | a. The patient is categorized wrongly. b. The specific treatment strategy is wrong. | Suboptimal treatment | Patient | 4 | 3 | 12 | Intolerable | The system shall be periodically updated. | 3 | 2 | 6 | Accepted | UD-CRS-09 UD-SRS-07
UD-H5 | Data availability | Unsuccessful treatment for specific patient due to lack of data | 1. The Udecide system recommends a treatment. 2. The patient recieves the treatment, but the health state deteriorates as a result. | Patient is exposed to a treatment that worsens the initial health condition. | Deterioration of symptoms | Patient | 1 | 5 | 5 | Broadly accepted | The risk is considered to be outweighted by the possible benefits of the treatment. | - | - | - | Accepted | UD-CRS-09 UD-SRS-07
UD-H6 | Patient identity and information | Malicious, unauthorized access to the Udecide system | 1. Someone gains unauthorized access to the Udecide system. 2. Patient information is stolen from the system | Patient's identity and health information can be widely shared and misused. | Identity and health information is misused. | Patient | 4 | 3 | 12 | Intolerable | a. Udecide is only available through a login-function. b. Conformity with ISO-82304. | 4 | 1 | 4 | Accepted | US-CRS-03 UD-SRS-23 MDR 2017:745 ISO-82304